XXX

Afonso Gusmão1,2,3,* ∙ Diogo F. Oliveira3,4.5 ∙ Frederico Castelo Ferreira1,2 ∙ Marco Leite3,* ∙

Jia An6∙ Arlindo Silva6,* ∙ U-Xuan Tan6,*

1 Department of Bioengineering and iBB—Institute for Bioengineering and Biosciences, Instituto Superior Técnico, Universidade de Lisboa, Av. Rovisco Pais, 1049-001 Lisbon, Portugal

2 Associate Laboratory i4HB—Institute for Health and Bioeconomy, Instituto Superior Técnico, Universidade de Lisboa, Av. Rovisco Pais, 1049-001 Lisbon, Portugal

3 IDMEC, Instituto Superior Técnico, Universidade de Lisboa, Av. Rovisco Pais, 1049-001, Lisbon, Portugal

4 Terraprima – Serviços Ambientais, Centro de Negócios do Porto Alto, Fração S, Avenida das Nações Unidas 97, 2135-199 Samora Correia, Portugal

5 MARETEC – Marine, Enviroment and Technology Centre, LARSyS, Instituto Superior Técnico, Universidade de Lisboa, Avenida Rovisco Pais 1, 1049-001 Lisboa, Portugal

6  Engineering Product Development, Singapore University of Technology and Design, 8 Somapah Road, Singapore 487372, Singapore

*Corresponding authors: afonsogusmao@tectico.ulisboa.pt, marcoleite@tecnico.ulisboa.pt, arlindo_silva@sutd.edu.sg, uxuan_tan@sutd.edu.sg

Abstract

XX

Keywords: XXX

Introduction

Paragraph 1 — Why cryogenic 3D bioprinting matters

Start broadly with the relevance of cryogenic 3D bioprinting and related freezing-based biofabrication strategies.
This first paragraph should establish why the area matters before talking about freezing-front dynamics.
Possible points:

- cryogenic bioprinting expands the processing window for soft and temperature-sensitive biomaterials;

- freezing can provide temporary shape stabilization during fabrication;

- cryogenic processing is also relevant because it can generate porous architectures and influence transport properties after freeze-drying or subsequent stabilization steps;

- more broadly, controlling freezing conditions is important for reproducibility, structural fidelity, and final scaffold performance.

The tone here should remain application-driven, not yet control-driven.

Paragraph 2 — Why freezing-front dynamics matter

After the broader cryobioprinting context, narrow to the core physical mechanism.
Key idea:

- final structure is not determined only by the minimum temperature or nominal cooling condition, but by how the freezing front propagates through the material over time.

This paragraph should connect freezing-front position and velocity to:

- local thermal gradients,

- ice-crystal growth conditions,

- and ultimately the spatial development of pore architecture.

Paragraph 3 — Limitation of fixed-temperature protocols

Then introduce the engineering limitation.
Key idea:

- in practice, freezing is often imposed through fixed plate temperatures or simple cooling protocols;

- these do not guarantee controlled or uniform front propagation;

- because thermal resistance and latent-heat effects evolve during the process, front velocity usually changes over time even under constant boundary conditions.

This is the paragraph where the problem becomes explicit.

Paragraph 4 — Gap in the literature

Now position the gap more sharply.
Suggested logic:

- many studies report empirical links between imposed freezing conditions and resulting structures;

- fewer studies validate predictive freezing models against embedded thermal data;

- fewer still use such validated models to design time-varying freezing trajectories;

- experimental deployment of such designed trajectories is especially rare.

Paragraph 5 — Proposed solution

This paragraph should introduce the paper’s solution framework.
Suggested content:

- a controllable cryostage provides the physical platform;

- a reduced-order model is used to represent executable plate-temperature dynamics;

- a predictive freezing model is validated against measured probe temperatures;

- an AI-assisted open-loop trajectory-design layer is used to generate executable reference-temperature profiles for front regulation.

Important: the cryostage should appear here as part of the proposed solution stack, not as the paper’s protagonist.
Any detailed cryostage characterization should be kept out of the main paper and moved to the Supplementary Material.

Paragraph 6 — Aim and contributions

Close the introduction with a compact aim/contribution paragraph.
It should state that the work addresses the engineering problem of regulating freezing-front dynamics through a hybrid experimental-computational framework and then list the main contributions in sentence form, not bullets.

A possible closing logic is:

- integration of a reduced-order cryostage response model into the control framework,

- validation of a predictive freezing model,

- formulation of an AI-assisted open-loop trajectory-design problem based on front-position tracking,

- and experimental demonstration of the designed trajectories.

Figure 1: Schematic of proposed approach (also Graphical Abstract)

In fixed-grid phase-change models, enthalpy-based formulations are commonly used to avoid explicit tracking of the solid–liquid interface, with the phase transition represented over a mushy region to reduce sharp numerical discontinuities [Dutil_2011.]

Experimental Setup

Mechanical and electrical design

The manufacturing phase of the LusoIceMaker was carried out at the Digital Gastronomy Lab of Singapore University of Technology and Design (Singapore), which is specialised in 3D food printing, where the structural, thermal, and electronic subsystems were assembled into a modular platform for directional freezing.  The system's adaptability was achieved by combining custom-designed mechanical parts with off-the-shelf thermoelectric, sensing, and control components. The structure of the system, shown in Figure X as a support for the cryo-stage, was designed using SolidWorks (Dassault Systèmes SolidWorks Corporation, Waltham, MA, U.S.) and fabricated by fused filament fabrication (FFF) using polylactic acid (PLA) on a Bambu Lab X1 Series (Shenzhen Tuozhu Technology Co., Ltd., Shenzhen, China). The main heat-transfer elements, namely the cold bed plate and the heat exchangers, as well as the chiller, thermoelectric modules (TEMs), sensors, microcontroller, cables, drivers, and associated electronic components, were obtained from third parties. Detailed information on the bill of materials (BOM), component sourcing, and assembly procedure is provided in Table S1.

Cryostage Specifications

The interface between the thermoelectric assembly and the sample mould consists of an aluminium plate measuring 100 × 100 × 8 mm. As shown in Figure X, the device uses four thermoelectric modules arranged in a 2 × 2 configuration. This cryostage represents a higher-power, faster-response version of our previously reported design [X], which relied on a cascaded two-stage thermoelectric architecture to maximise the achievable ΔT. The selected modules were HiTemp ETX7-16-F1-4040-TA-W8 units (Laird Thermal Systems, NC, USA), each with a maximum cooling capacity (Qc,max) of 84.1 W at ΔT=0 and a maximum temperature differential (ΔTmax) of 83.2°C at Qc=0. On the hot side, the thermoelectric stack was coupled to a liquid-cooling sink consisting of an aluminium block measuring 80 × 80 × 15 mm with 8 mm-diameter internal channels, through which coolant was circulated by a DC-0506 recirculating thermostatic bath (SanLiang, China). This unit has a rated power of 800 W, allowing continuous heat removal during prolonged subzero operation. To reduce interfacial thermal resistance, Kryonaut thermal paste (Thermal Grizzly, Lippstadt, Germany) with a thermal conductivity of 12.5 W·m·-1·K−1 was applied between the contacting surfaces.

A custom clamping structure was designed to align the cold plate, thermoelectric modules, and liquid-cooling sink into a single modular unit. In addition to ensuring mechanical stability, this structure provides thermal insulation due to PLA's low thermal conductivity and facilitates cable routing and fluidic integration. This modular arrangement allows the cryostage to be installed or removed without permanently modifying the printing platform, while also simplifying maintenance of the cooling stack, sensor wiring, and fluidic connections.

Thermal control and sensing

The electrical architecture of the cryostage is divided into three main parts: temperature sensing, power actuation, and control. Plate temperature is monitored using a three-wire PT100 RTD inserted into a dedicated cavity in the aluminium cold plate, as shown in Figure X. The RTD signal is acquired through a MAX31865 interface board (Adafruit Industries LLC, Brooklyn, NY, USA) and processed by an ESP32 microcontroller (Espressif Systems, Shanghai, China), which executes the control algorithm and manages communication with the user interface. Power delivery to the thermoelectric modules is achieved through BTS7960 driver modules supplied by two 24 V, 25 A power supplies, with each power supply feeding two Peltier TEMs, as shown in the circuit diagram in Figure SX. The control firmware is implemented in the Arduino framework using C++, and a PID controller based on the Arduino PID library is used to regulate the cold plate temperature from the RTD feedback. The controller operates with fixed-temperature setpoints and time-varying reference trajectories, thereby enabling baseline freezing experiments and programmed cooling protocols. The thermocouple signals were acquired through MAX31856 thermocouple interface boards, enabling simultaneous acquisition of cold-plate, in-sample, and surrounding-air temperature histories.

Mould Design

The mould used in the freezing experiments is fabricated by fused filament fabrication (FFF) using polylactic acid (PLA) on the same Bambu Lab X1 Series printer described above. The mould has an overall height of approximately 17 mm, an inner diameter of 15 mm, and a wall thickness of 2 mm. Three lateral holes are incorporated at the nominal measurement heights to allow insertion of the thermocouples. In addition, an internal ledge is included at a height of 15 mm to define the filling level and improve reproducibility, with water added to this mark in all tests. Type T thermocouples were inserted through the side cavities and fixed externally using UV-curable epoxy. After insertion and fixation, the effective thermocouple junction heights were measured, yielding real measurement positions of 3.0, 6.2, and 11.0 mm relative to the cold-plate surface. The junctions were located approximately 0.5 mm from the inner mould wall, corresponding to rprobe=7.0 mm.

As shown in Figure X, an additional PLA mould holder was fabricated to position the mould at the centre of the cold plate and ensure consistent placement during the experiments. The remaining three thermocouples were located within the sample at the effective heights listed above, enabling the water sample's thermal history to be recorded during freezing.

An additional external thermocouple was positioned outside the mould, adjacent to its lateral wall, at approximately mid-height (∼8 mm). Its purpose was not to measure the bulk laboratory room temperature, but rather to monitor the local air temperature in the immediate vicinity of the mould during freezing. The cryostage clamping structure includes lateral walls surrounding the working area that help retain the cold air generated near the cold bed and reduce its rapid dissipation via natural convection. Accordingly, this measurement was interpreted as the local surrounding-air temperature around the mould.

Interface and sampling

To support control and experimental monitoring, a custom desktop interface is implemented in Python and connected to the ESP32 through serial communication. This interface allows the user to define and apply temperature setpoints or programmed trajectories, while simultaneously visualising and recording, in real time, the temperatures measured by the RTD and the thermocouples embedded in the mould. In the implemented firmware, sensor acquisition, control update, and telemetry streaming are performed at 100 ms intervals. In Figure SX interface is shown, including the control buttons on the right-hand side and the real-time telemetry from the different sensors on the left-hand side.

Experimental protocol

Sensor’s calibration

The temperature sensors were calibrated by a two-point linear regression using reference temperatures measured inside a laboratory temperature- and humidity-controlled chamber [manufacturer, blbla). Calibration was performed at 5 °C and 30 °C. For each sensor channel, the firmware implemented a linear correction of the form:

|  | Tcal=aTraw+b | (1) |
| --- | --- | --- |

Where Traw is the raw sensor reading, and Tcal is the corrected temperature used for data acquisition and subsequent analysis. The fitted coefficients aand b for the cold-plate RTD, the three in-sample thermocouples, and the external mould-side thermocouple are presented in Table S1.

Cryostage thermal response characterisation

Dry cooling assays were performed to characterise the cryostage thermal response in the absence of a sample. The plate temperature was set to constant setpoints of -5 °C, -10 °C, -15 °C, and -20 °C, with three independent runs (n=3) for each condition. During these assays, the calibrated plate temperature was continuously recorded to assess transient response and repeatability. The calibrated plate-temperature telemetry was recorded at a median sampling interval of 0.111. These dry-cooling traces were subsequently used to estimate the empirical plate-response model used in the open-loop trajectory-design workflow (Section 4.1).

Water directional freezing characterization

Freezing experiments were performed using purified water [purification system brand, model, country] as a model material. For each experiment, the PLA mould described in Section 2.1.3 was positioned at the centre of the cold plate and pre-cooled at the target setpoint before filling. Water was added only after the thermocouple readings had reached a stable plateau. The filling event was defined as the start of the freezing assay. Freezing experiments were then carried out at constant cold-bed temperatures of -10 °C, -15 °C, and -20 °C, with five independent replicates per condition (n=5). The cold-plate RTD and all thermocouple signals were recorded throughout each run.

Freezing simulation model

Model domain, geometry, and boundary conditions

The freezing problem was represented as a two-dimensional axisymmetric thermal problem in the (r,z) plane. The computational domain was defined as the rectangle

|  | Ω= r,z : 0≤r ≤Rout, 0 ≤z ≤Htotal | (2) |
| --- | --- | --- |

where Rout=Rin+twall. In the implemented calibration geometry, the inner cavity radius was Rin=7.5 mm, the mould wall thickness was twal=2 mm, the fill height was Hfill=15 mm, and the total mould height was  Htotal=17 mm.

The domain was partitioned into three physical regions: the mould, the lower fillable cavity, and the upper headspace. Using the axisymmetric coordinates, these regions were defined as

|  | Ωfill= r,z : 0≤r ≤Rin, 0 ≤z ≤Hfill, | (3) |
| --- | --- | --- |
|  | Ωhead= r,z : 0≤r ≤Rin, Hfill ≤z ≤Htotal, | (4) |

and

|  | Ωmould=Ω \ (Ωfill ⋃ Ωhead ) | (5) |
| --- | --- | --- |

Before filling, both the fillable cavity and the headspace were treated as air, while the mould region was assigned PLA properties. After filling, only Ωfill was replaced by water, whereas Ωhead  remained air throughout the simulation. Thus, the model explicitly includes the internal air headspace above the liquid level. In contrast, the external surrounding air around the mould was not represented as a separate domain, but through effective convective boundary conditions.

The thermal field T(r,z,t) was driven by the imposed plate temperature at the lower boundary and by convective losses at the exposed boundaries. The bottom boundary,

|  | Γb​={(r,0):0≤r≤Rout​}, | (6) |
| --- | --- | --- |

was specified as a Dirichlet condition,

|  | T(r,0,t)=Tplate​(t) | (7) |
| --- | --- | --- |

In the constant-temperature validation runs, Tplate(t) reduced to a fixed plate setpoint. The top boundary,

|  | Γt​={(r,Htotal​):0≤r≤Rout​} | (8) |
| --- | --- | --- |

and the outer lateral boundary,

|  | Γo​={(Rout​,z):0≤z≤Htotal​} | (9) |
| --- | --- | --- |

were modelled using Robin boundary conditions,

|  | -k∇T⋅n=h​T-T∞​ on Γt ⋃ Γo​, | (10) |
| --- | --- | --- |

where h=2 Wm-2K-1 and T∞is the effective local surrounding-air temperature defined in Section 2.1.3. The boundary at r=0 was treated as the axis of symmetry, corresponding to zero radial heat flux. For each constant-plate validation condition, T∞ was calculated as the mean external mould-side thermocouple temperature across the corresponding freezing assays. The calibration values are reported in Table S2 and were used to define T∞Tplate by linear interpolation, with nearest-value clamping outside the measured range.

The model also included an explicit pre-fill stage. In the steady pre-fill mode implemented in the solver, the empty mould problem was first solved as a steady conduction problem with the cavity and headspace both treated as air:

|  | ∇⋅k∇T=0​ in Ω | (12) |
| --- | --- | --- |

subject to the same bottom Dirichlet and top/side Robin boundary conditions described above. This produced the pre-cooled temperature field used as the initial state immediately before filling.

The fill event was then imposed numerically at time t=tfill by resetting only the lower cavity region to the fill temperature Tfill,

|  | Tr,z, tfill+=Tfill, &(r,z)∈ ΩfillT(r,z,tfill-), &(r,z)∈ Ω / Ωfill ,​ | (13) |
| --- | --- | --- |

so that the mould and headspace retained their pre-cooled state. In the implemented code, this overwrite was applied only to the lower water-filled region, excluding the shared interface degrees of freedom at the mould wall.

Freezing model formulation

The freezing of the cylindrical samples was treated as a Stefan-type phase-change problem. However, the solid-liquid interface was not tracked explicitly as a moving boundary. Instead, the thermal problem was solved on a fixed axisymmetric domain using a continuum enthalpy-family formulation [1], [2], implemented here through a temperature-based apparent heat-capacity regularization [3], [4]. Natural convection in the liquid phase was not explicitly resolved. Instead, the model adopted a conduction-dominated formulation focused on the transient temperature field and freezing-front progression. This simplification is consistent with prior phase-change studies showing that conduction-only models can be adequate when convective transport remains limited [5].  The model was designed to predict the transient temperature field and freezing-front progression under prescribed thermal boundary inputs, while remaining computationally efficient for trajectory design rather than providing a full thermo-fluid description of the liquid phase.

Within each subdomain, the model solved transient heat conduction in axisymmetric form. In the mould cavity region, after the filling event, the governing equation can be written as:

|  | ρ(T)cc,effT∂T∂t=1r ∂∂rrkT∂T∂r+ ∂∂zkT∂T∂z | (15) |
| --- | --- | --- |

where Tr, z, t is the temperature field, k(T)is the thermal conductivity, ρ(T)is the density, and cp,eff(T) the apparent heat capacity used in the transient storage term. The left-hand side represents transient energy storage, while the right-hand side represents radial and axial heat conduction in cylindrical coordinates.

The mould and headspace domains were assigned constant properties for PLA and air, respectively (Table S3). In the water-filled cavity, liquid-water and ice thermal conductivity k(T), specific heat capacity cp(T), and density ρ(T) were assigned according to the local phase state, defined by the sign of T-Tf, where Tf=0ºC is the reference freezing temperature of water. The temperature dependent properties were linearly interpolated from tabulated values reported by Incropera et al. [6] and summarized in Table S4. Phase selection followed a strict threshold: ice properties were used for T<Tf, whereas liquid-water properties were used for T≥Tf. Thus, at T=Tf, the base properties corresponded to liquid water, with latent heat added separately through Eq. (16). Liquid-water density was fixed at ρ=1000 kg m-3, while ice density was interpolated from the temperature-dependent values in Table S4. This retained the liquid–solid density contrast while neglecting intra-phase liquid density variations, consistent with phase-change formulations where density changes are simplified unless volume change or thermo-mechanical coupling is explicitly modelled [7].

The latent heat contribution was incorporated through cp,eff(T), term rather than introduced through explicit front tracking, which allows the same temperature equation to be solved across both unfrozen and frozen states [3], [4]. For this purpose, a liquid-fraction function, fl(T), was used to regularize the transition between liquid and solid states. The apparent heat capacity was expressed as:

|  | cp,eff(T)=cpT+LdfldT | (16) |
| --- | --- | --- |

where cp(T) is the temperature-dependent heat capacity presented before and L=334 kJ kg-1 is the latent heat of fusion of water. The latent term was distributed uniformly over a finite temperature interval ΔTmushy=0.5 centred at the reference freezing temperature, Tf. This interval was used as a numerical superficial mushy region to regularise the discontinuity associated with the phase transition, rather than as an explicitly resolved physical mushy zone in pure water [8], [9]:

|  | cp,lat(T)=LlatentΔTmushy, Tf-ΔTmushy2 ≤T≤Tf+ΔTmushy20, &otherwise | (17) |
| --- | --- | --- |

Thus, cp,eff(T)=cp(T)+cp,lat(T), and the same temperature-based governing equation was solved across both unfrozen and frozen states. The phase transition was therefore represented as a regularised diffuse thermal interval rather than as an explicitly tracked sharp interface.

Numerical implementation and extraction of observables

Finite-element implementation and time integration

The governing equations were solved using a finite-element formulation implemented in FEniCSx. The weak form was defined using the Unified Form Language (UFL), allowing the axisymmetric weighting, material-dependent coefficients, and Robin boundary terms to be expressed directly in variational form [10], [11], [12]. FEniCSx was then used to assemble the finite-element systems, which were solved through the Portable, Extensible Toolkit for Scientific Computation (PETSc) [13]. The resulting linear systems were solved using a conjugate-gradient Krylov solver with Hypre preconditioning [14]. The computational mesh was generated as a triangular discretisation of the axisymmetric (r,z) domain defined in Section 3.1. The temperature field was approximated using first-order continuous Lagrange elements. Axisymmetry was introduced at the weak-form level by weighting all domain and boundary integrals by the radial coordinate r.

The transient problem was time-integrated using an implicit Euler scheme. At each time step (Δt), keff  and ρcp,eff were evaluated from the previous temperature field and kept fixed while solving for the next temperature field. This avoided an iterative nonlinear solve while retaining temperature-dependent material updates and the apparent heat-capacity term. The relative and absolute tolerances were set to 10-10and 10-12, respectively. The steady pre-fill problem used the same finite-element formulation with tighter tolerances, since it defined the initial thermal state before filling.

Numerical sensitivity

Mesh and time-step sensitivity were assessed against finer reference simulations by varying the number of finite-element mesh divisions in the radial direction, Nr, the number of finite-element mesh divisions in the axial direction, Nz, and the transient time-step size, Δt. Probe-temperature histories, front-position histories, front-passage times, and freezing-completion times were used as comparison metrics. The selected values, Nr=180, Nz=408, and Δt=0.25 s, were used because further refinement produced only minor changes in these observables (Table S5).

Front-extraction parameters were treated separately from the finite-element discretisation. The computed finite-element temperature field was evaluated on independent post-processing sampling lines or grids, which did not alter the numerical solution. For the centreline and near-wall front-position trajectories, Nz,front=800 axial sampling points were used. When curved-front extraction was enabled, Nr,front=25 radial positions and Nz,front,curve=400 axial sampling points were used to reconstruct zf(r,t). A separate sampling-sensitivity check confirmed that these values were sufficient for stable front extraction (Table S5).

Extraction of observables and exported outputs

Probe-temperature histories were extracted by sampling the finite-element temperature field at the thermocouple positions defined in Section 2.1.3. These corresponded to z=3.0, 6.2, and 11.0mm at the measured near-wall radial location. The extracted histories, T3.0(t), T6.2(t), and T11.0(t), were exported with the absolute simulation time and the time elapsed after filling.

Freezing-front position was extracted as a post-processed isotherm from the computed temperature field. Unless otherwise stated, the front criterion was Tfront=Tf. At each recorded front-output time, T(r,z,t) was sampled along vertical lines in Ωfill, defined in Eq. (3) and the front height zft was obtained by linear interpolation at the first crossing of Tfront from the cold plate upwards. Lines entirely warmer than Tfront were assigned zf=0, whereas lines entirely colder than Tfrontwere assigned zf=Hfill. Routine front tracking produced two axial trajectories: a centreline front, zf,c(t), and a near-wall front, zf,w(t), evaluated at the thermocouple radial coordinate defined in Section 2.1.3. For curved-front analysis, the same isotherm-crossing procedure was applied across multiple radial positions to reconstruct zf(r,t). Accordingly, zfwas treated as a threshold-based observable, not as an explicitly tracked interface.

Freezing was considered complete at the first recorded front-output time for which

|  | max⁡(r,z)∈ΩfillT(r,z,t)≤Tfreeze, | (18) |
| --- | --- | --- |

where Tfreeze=Tf-(ΔTmushy/2+ΔTsubcooling). In the simulations reported here, ΔTsubcooling=0, so Tfreeze corresponded to the lower bound of the apparent heat-capacity interval defined in Eq. (17). This domain-level criterion avoided defining complete freezing from a single probe location or front-sampling line.

Front-passage times were obtained from the post-processed front trajectory defined above. For a selected trace zf(t), the passage time at an axial position ziwas defined as

|  | tpass(zi)=min⁡{t:zf(t)≥zi}. | (19) |
| --- | --- | --- |

Passage times were extracted at the thermocouple axial positions defined in Section 2.1.3. When ziwas crossed between two recorded front-output times, tpass(zi)was obtained by linear interpolation of the neighbouring zf(t)values.

Front velocities were derived from the post-processed front-position trajectories, rather than solved as primary variables. For each selected front position zf(t), the axial velocity was estimated from its temporal variation. For comparisons between thermocouple heights, segment-wise average velocities were calculated from the passage times defined in Eq. (19):

|  | vi→j=zj-zitpass(zj)-tpass(zi) | (20) |
| --- | --- | --- |

This provided an effective front-propagation velocity between experimentally instrumented axial positions, without relying on pointwise differentiation of zf(t).

Simulation outputs were exported in two formats. Full temperature fields were written in XDMF/HDF5 format and visualised in ParaView 6.0.1 (Kitware Inc., Clifton Park, NY, USA) for field-level inspection. Post-processed observables were exported as CSV files for validation, plotting, and trajectory-design workflows. These included the probe-temperature histories, the centreline and near-wall front-position trajectories and associated front metrics, and the curved-front profiles zf(r,t)when radial front reconstruction was enabled.

Open-loop trajectory design and optimization

Control objective and front-position tracking

The trajectory-design problem was formulated as a model-based open-loop input-design problem with respect to freezing-front regulation: a finite-dimensional parameter vector θ defines a reference-temperature trajectory Tref(t), which is evaluated through the cryostage response model and freezing solver. The trajectory is specified as a function of time and is not updated during each run using feedback from the freezing-front position. In the computational workflow, Tref(t) is mapped to a modelled plate-temperature trajectory Tplate(t), which is then imposed as the lower boundary condition in the freezing solver. The resulting temperature field T(r,z,t) is post-processed to obtain the centreline freezing-front position zf(t), which is used to compute the scalar objective J(θ). This cascade is summarized in Figure X. The trajectory parameterization used to construct Tref(t) from θ is described in Section 4.2.

During optimization, the mapping from the reference-temperature trajectory to the plate-temperature boundary condition was represented by an empirical setpoint-dependent first-order response model obtained from the dry cryostage characterization assays described in Section 2.2.2. For a constant Tref, the plate temperature, the model assumes that the plate temperature relaxes toward Tplate,ss(Tref), which is the fitted asymptotic plate temperature associated with that reference, with setpoint-dependent response time τ(Tref). The fitted values of Tplate,ss(Tref) and τ(Tref) are reported in Table SX, and the corresponding dry-cooling curves with first-order model overlays are shown in Figure SX. The continuous-time form of the model is

|  | dTplatedt=Tplate,ssTref-Tplateτ(Tref) | (21) |
| --- | --- | --- |

In this discrete implementation used in the trajectory-design workflow, this response was evaluated as

|  | Tplate,i​=αi​Tplate,i-1+1-αi,Tplate,ssTref,i-1, | (22) |
| --- | --- | --- |
|  | αi=exp-Δtiτ(Tref,i-1) |  |

where Δti=ti-ti-1. Between the characterized setpoints, Tplate,ss(Tref) and τ(Tref) were obtained by interpolation of the fitted lookup values. This model represents the observed dynamic plate-temperature response to a requested reference-temperature trajectory. Similar input-design formulations have been used in Stefan and solidification-control studies, most directly in boundary-cooling design for desired freezing-interface motion [15].

The trajectory-design objective was defined using front-position tracking rather than direct instantaneous front-velocity tracking. As described in Section 3.3, the freezing front was obtained as a post-processed observable from the simulated temperature field. Front velocity would therefore require numerical differentiation of the extracted front-position trajectory. Because numerical differentiation of sampled data is ill-posed and can amplify noise, making derivative estimates sensitive to sampling and smoothing choices [16], , front velocity was treated as a diagnostic quantity rather than as the primary optimization target.

The target freezing behaviour was defined as an approximately linear advance of the centreline front through the filled sample height. The reference front trajectory was defined as

|  | zf,reft=0, ,t<t0, Hfillt-t0tend-t0, t0≤ t≤tend Hfill, t>tend | (23) |
| --- | --- | --- |

where Hfill is the filled sample height, t0 is the start of the objective-evaluation window, and tend is the design horizon. This reference corresponds to an approximately constant average front velocity over the active interval,

|  | vref=Hfilltend-t0 | (24) |
| --- | --- | --- |

For each candidate trajectory, the simulated centreline front position was compared with the reference trajectory at the valid post-fill sample times. The main tracking term was defined as

|  | Jtrack​θ=1Nki=1Ntzfti;θ-zf,reftiHfill2, | (25) |
| --- | --- | --- |

where Nt is the number of time samples included in the objective calculation. Normalization by Hfill makes the tracking error dimensionless and expresses deviations relative to the filled sample height.

|  | Jsmooth​θ=1Nk-1j=1Nk-1(Tj+1-Tj)/(tj+1-tj)ΔTref/Δtref | (26) |
| --- | --- | --- |

where Tj is the reference temperature at knot time tj, Nk is the number of knots, ΔTref is the allowed reference-temperature range, and Δtref is the trajectory duration.

The scalar objective minimized during trajectory design was written as

|  | Jθ=wtrackJtrackθ+wsmoothJsmoothθ | (27) |
| --- | --- | --- |

where wtrack, wsmooth, and wcomp are weighting factors. In the active linear full-process formulation, Jtrack was the primary term, while Jsmooth acted as a secondary regulariser on the requested reference-temperature trajectory. The trajectory parameterization and admissibility constraints imposed on Tref(t) are described in Sections 4.2 and 4.3.

Trajectory parameterization and admissible reference inputs

The Tref trajectory was parameterized by a finite set of temperature values prescribed at fixed knot times. For a trajectory with N knots, the design vector was defined as θ=(θ1,θ2,…,θN), where θi is the requested reference temperature at the fixed time ti. The knot times were selected before trajectory evaluation and were not optimized, 0=t1<t2<…<tN=tend. Between consecutive knot times, Treft;θ was obtained by linear interpolation,

|  | Treft;θ=θi+θi+1-θiti+1-tit-ti, | (28) |
| --- | --- | --- |
|  | ti≤t≤ti+1 |  |

For times outside the knot interval, the first or last knot temperature was held constant. Thus, θ defined the temperature values of the requested trajectory, while the time support defined the temporal resolution of the trajectory representation. Different knot schedules can be used within the same formulation; the specific schedule and number of knots used in each reported case are given with the optimization settings.

Candidate trajectories were constrained to remain within the allowable reference-temperature range,

|  | Tmin≤θi≤Tmax, | (29) |
| --- | --- | --- |
|  | i=1,…,N |  |

In the active workflow, the requested trajectory was also restricted to be monotone non-increasing, θ1≥θ2≥…≥θN. This constraint allows the requested cryostage reference to remain constant or become colder during freezing, while excluding warming segments that were not supported by the characterized operating regime. After these representation-level checks, the admissible Tref(t) was passed to the empirical cryostage-response model to obtain Tplate(t), which was then used as the lower boundary condition in the freezing solver.

Results and Discussion

Validation experiments under constant temperature protocols

References

[1]	V. R. Voller, C. R. Swaminathan, and B. G. Thomas, ‘Fixed grid techniques for phase change problems: A review’, Numerical Meth Engineering, vol. 30, no. 4, pp. 875–898, Sep. 1990, doi: 10.1002/nme.1620300419.

[2]	C. R. SWAMINATHAN and V. R. VOLLER, ‘ON THE ENTHALPY METHOD’, International Journal of Numerical Methods for Heat & Fluid Flow, vol. 3, no. 3, pp. 233–244, Mar. 1993, doi: 10.1108/eb017528.

[3]	D. Poirier and M. Salcudean, ‘On Numerical Methods Used in Mathematical Modeling of Phase Change in Liquid Metals’, J. Heat Transfer, vol. 110, no. 3, pp. 562–570, Aug. 1988, doi: 10.1115/1.3250529.

[4]	S. R. Runnels and G. F. Carey, ‘Finite Element Simulation of Phase Change Using Capacitance Methods’, Numerical Heat Transfer, Part B: Fundamentals, vol. 19, no. 1, pp. 13–30, Jan. 1991, doi: 10.1080/10407799108944954.

[5]	V. R. Voller and I. Vus̆anović, ‘A criterion for determining when convection needs to be considered in calculations of solidification/melting in thermal cavities’, International Communications in Heat and Mass Transfer, vol. 159, p. 108301, Dec. 2024, doi: 10.1016/j.icheatmasstransfer.2024.108301.

[6]	F. P. Incropera and D. P. Dewitt, Eds, Fundamentals of heat and mass transfer, 7th ed. Hoboken, NJ: John Wiley, 2007.

[7]	J. Dallaire and L. Gosselin, ‘Various ways to take into account density change in solid–liquid phase change models: Formulation and consequences’, International Journal of Heat and Mass Transfer, vol. 103, pp. 672–683, Dec. 2016, doi: 10.1016/j.ijheatmasstransfer.2016.07.045.

[8]	M. Bhattacharya, T. Basak, and K. G. Ayappa, ‘A fixed-grid finite element based enthalpy formulation for generalized phase change problems: role of superficial mushy region’, International Journal of Heat and Mass Transfer, vol. 45, no. 24, pp. 4881–4898, Nov. 2002, doi: 10.1016/S0017-9310(02)00178-3.

[9]	J. Malík and O. Souček, ‘A semi-analytical benchmark for the Stefan problem in arbitrary dimension – assessing accuracy of enthalpy-based methods’, International Journal of Numerical Methods for Heat & Fluid Flow, vol. 32, no. 11, pp. 3598–3635, May 2022, doi: 10.1108/HFF-09-2021-0647.

[10]	M. S. Alnæs, A. Logg, K. B. Ølgaard, M. E. Rognes, and G. N. Wells, ‘Unified form language: A domain-specific language for weak formulations of partial differential equations’, ACM Trans. Math. Softw., vol. 40, no. 2, p. 9:1-9:37, Mar. 2014, doi: 10.1145/2566630.

[11]	M. Alnæs et al., ‘The FEniCS Project Version 1.5’, Archive of Numerical Software, vol. 3, no. 100, Dec. 2015, doi: 10.11588/ans.2015.100.20553.

[12]	I. A. Baratta et al., ‘DOLFINx: The next generation FEniCS problem solving environment’, Dec. 31, 2025, Zenodo. doi: 10.5281/zenodo.18101307.

[13]	S. Balay, S. Abhyankar, and M. F. Adams, ‘PETSc/TAO Users Manual’, 2024.

[14]	R. D. Falgout and U. M. Yang, ‘hypre: A Library of High Performance Preconditioners’, in Computational Science — ICCS 2002, P. M. A. Sloot, A. G. Hoekstra, C. J. K. Tan, and J. J. Dongarra, Eds, Berlin, Heidelberg: Springer, 2002, pp. 632–641. doi: 10.1007/3-540-47789-6_66.

[15]	S. Kang and N. Zabaras, ‘Control of the freezing interface motion in two-dimensional solidification processes using the adjoint method’, International Journal for Numerical Methods in Engineering, vol. 38, no. 1, pp. 63–80, 1995, doi: 10.1002/nme.1620380105.

[16]	F. Van Breugel, J. N. Kutz, and B. W. Brunton, ‘Numerical Differentiation of Noisy Data: A Unifying Multi-Objective Optimization Framework’, IEEE Access, vol. 8, pp. 196865–196877, 2020, doi: 10.1109/ACCESS.2020.3034077.

[17]	‘Marlin Firmware, v2.0, https://marlinfw.org/, 2024 (accessed 20 November 2024)’.